СОП № ЛММБ 2-7-2017-09

Идентификация бактериального штамма по биохимическим свойствам с использованием анализатора GenIII OmniLog

- Введение, цель

Настоящая методика устанавливает порядок идентификации микроорганизма, депонированного в КЭМТК ИХБФМ, по его биохимическим свойствам.

- Назначение

Идентификация микроорганизма является необходимым этапом при депонировании микроорганизма в Коллекции ЭМТК

- Термины и определения

СОП – стандартная операционная процедура;

Асептика – комплекс мер направленных на предупреждение попадания в рабочую зону сторонних микроорганизмов;

Селективные среды – питательные среды для выделения определенных микроорганизмов за счет создания благоприятных для них условий роста и неблагоприятных условий для сопутствующих микроорганизмов других видов;

РПА – рыбо-пептонный агар;

СМА – сердечно-мозговой агар;

СМБ – сердечно-мозговой бульон.

- Пересмотр

Данная СОП вводится впервые.

- Материалы и оборудование

  - Материалы и реактивы

| Наименование основных реактивов и материалов | НТД, производитель, страна |
| --- | --- |
| РПА | ТУ 9385-012-14237183-07 |
| СМА | BioMerieux, Франция |
| Микропланшет GEN III MicroPlate | BioLog, США |
| Инокулятор Inoculatorz | BioLog, США |
| Наконечники 1250 мкл для инокулирования микропланшетов | BioLog, США |
| Инокулирующая жидкость IF-A | BioLog, США |
| Инокулирующая жидкость IF-B | BioLog, США |
| Инокулирующая жидкость IF-C | BioLog, США |
| Электронная пипетка Ovation | BioLog, США |
| Комплект автоматических пипеток с переменным объемом (4 шт) | «Ленпипет», Россия |
| Резервуар стерильный для использования с многоканальными пипетками | BioLog, США |
| Спиртовка лабораторная | ГОСТ 23932-90Е |
| Колбы мерные вместимостью 250 мл | ГОСТ 1770-74 |
| Флаконы градуированные c завинчивающейся крышкой, 500 мл | Isolab, Германия |
| Цилиндры мерные вместимостью 50 мл, 500 мл и 1 л | ГОСТ 1770-74 |
| Пробирки типа Eppendorf, 1.5 мл | Thermo, Россия |
| Наконечники для автоматических дозаторов объемом 200 мкл | «Eppendorf», США |
| Спирт этиловый, ректифицированный | ЛРС 000279/10 |
| Перекись водорода, медицинская | ГОСТ 177-88 |
| Вода дистиллированная рН от 5,0 до 7,0. | ГОСТ 6709-72 |
| Натрий хлористый, хч | ГОСТ 4233-77 |
| Чашки Петри, пластиковые диаметр 90 мм | Greiner Bio-One, Австрия |
| Бактериологическая петля пластиковая одноразовая | Citotest, Китай |

  - Оборудование

| Оборудование | НТД, производитель, страна |
| --- | --- |
| Биохимический прибор-анализатор GenIII OmniLog Plus | BioLog, США |
| Бокс биологической (микробиологической) безопасности II класс | Lamsystems, Россия |
| Турбидиметр | BioLog, США |
| Холодильник | Indesit, Италия |
| Весы электронные аналитические | Ohaus, США |
| Термостат суховоздушный лабораторный ТСвЛ-80 | ТУ-9452-006-07505566-2006 |
| Автоклав ВК-75 | ТЗМОИ, Россия |
| Вортекс | BioSan, Латвия |

  - Комплект спецодежды

| Одежда | НТД, производитель, страна |
| --- | --- |
| Колпак медицинский | ГОСТ 2313478 |
| Перчатки хирургические резиновые | ГОСТ 3-88 |
| Маска медицинская | ГОСТ EN 13795-1-2011 |
| Халат медицинский | ГОСТ 24760-81 |

- Помещения

Проведение работ осуществляется в боксовых помещениях, в которых находятся боксы биологической (микробиологической) безопасности II класса, а также в помещениях, предназначенных для общих молекулярно-биологических работ.

Процедура

  - Подготовительный этап

7.1.1. Подготовка персонала к проведению работ

– надеть медицинский халат и перчатки

7.1.2. Приготовление дезинфицирующего раствора

– приготовить 3% раствор перекиси водорода, в стеклянный цилиндр налить (100±1) мл 30% перекиси водорода и довести объем до 1000 мл водопроводной водой, данный раствор может быть использован в течение 48 ч;

7.1.3. Подготовка боксового помещения к работе

– обработать 3% раствором перекиси водорода поверхности помещения и оборудования до начала работ;

– обработать ламинарный бокс и помещение ультрафиолетовыми лучами до начала работ в течение 15 мин.

7.1.4. Приготовление 70% раствора этилового спирта

– налить в стеклянный цилиндр (70±1) мл 96% этилового спирта и довести объем до 100 мл дистиллированной водой;

7.1.5. Подготовка чашек Петри с питательной средой

– приготовить питательную среду согласно инструкции производителя;

– разлить питательную среду после автоклавирования в стерильные чашки Петри толщиной (4,00,5) мм и оставить для застывания при комнатной температуре.

7.1.6. Пересев культуры на неселективные питательные среды

– бактериологической петлей отобрать небольшое количество биомассы с поверхности колоний и высеять методом истощающего штриха на чашку Петри с питательной средой РПА или СМА. Инкубировать чашку согласно условиям, записанным в журнале депонирования микроорганизмов в течение 24 – 72 часов. Использовать выросшие колонии для дальнейшей идентификации культуры.

7.1.7. За час до начала засева микропланшетов включить биохимический прибор-анализатор GenIII OmniLog Plus для прогрева термостата-инкубатора.

7.2. Основной этап

7.2.1. Подготовка персонала к проведению работ

– надеть боксовый халат, перчатки, шапочку и медицинскую маску. Дальнейшую работу вести в условиях ламинарного бокса биобезопасности.

7.2.2. Засев микропланшетов GEN III MicroPlate и инкубирование планшетов в биохимическом анализаторе GenIII OmniLog Plus

–  выбрать тип инокулирующей жидкости (IF-A, IF-B или IF-C) согласно инструкции к прибору GenIII OmniLog Plus (Приложение) и предварительным данным о возможной таксономической принадлежности штамма, полученным на основе культуральных и морфологических свойств колоний и клеток исследуемого штамма;

– коснуться инокулятором выросшей колонии на чашке Петри, после чего поместить инокулятор в пробирку с инокулирующей жидкостью, опустить его до дна пробирки и растереть бактериальную массу по дну пробирки;

–  перемешать жидкость в пробирке с использованием вортекса и замерить оптическую плотность инокулированной культуры с использованием турбидиметра. Оптическая плотность должна быть не более 0.05 OD. Если оптическая плотность выше – разбавить полученную культуру дополнительной стерильной инокулирующей жидкостью;

– перелить инокулят в стерильный резервуар и из него электронной многоканальной пипеткой Ovation раскапать по 100 мкл во все лунки микропланшета GEN III MicroPlate;

– подписать микропланшет, закрыть крышкой и перенести в термостат-инкубатор биохимического анализатора GenIII OmniLog Plus;

–  включить программное обеспечение Microbial Identification Systems Software, внести запись о микропланшете в базу данных Microbial Identification Systems Software и выбрать протокол инкубирования, соответствующий инокулирующей жидкости (А, B, или C);

–  инкубировать планшет в термостате-инкубаторе при 33 °С в режиме с периодическим автоматическим считыванием данных. Максимальный срок инкубации составляет 36 часов, в течение этого времени прибор должен будет идентифицировать штамм;

–  распечатать лист идентификации, сложить в папку «Идентификация штаммов  GenIII OmniLog Plus» и внести полученные данные в электронную базу данных микроорганизмов КЭМТК;

– в случае, если таксономическую принадлежность штамма определить не удалось, засеять новый микропланшет GenIII OmniLog Plus той же культурой, взяв инокулирующую жидкость другого типа и поменяв протокол инкубирования;

– в случае выявления принадлежности штамма при повторном инкубировании, распечатать лист идентификации, сложить в папку «Идентификация штаммов  GenIII OmniLog Plus» и внести полученные данные в электронную базу данных микроорганизмов КЭМТК;

– если таксономическая принадлежность при повторном инкубировании не выявлена, идентифицировать микроорганизм согласно СОП № ЛММБ-2-4-2017-09 или СОП № ЛММБ-2-6-2017-09.

7.3. Завершающий этап

– замочить инкубированные микропланшеты в 6% растворе перекиси водорода, избегая образования воздушных пробок. Экспозиция не менее 6 ч.

– обработать 3% раствором перекиси водорода поверхности помещения и оборудования после окончания работ;

– обработать ламинарный бокс и помещение ультрафиолетовыми лучами после окончания работ в течение 15 мин.

- Охрана труда и техника безопасности

При проведении процедуры необходимо соблюдать следующие инструкции по технике безопасности и инструкции по биобезопасности:

- ИОТ – 02 Инструкция по ОТ для неэлектротехнического персонала по электробезопасности на I квалификационную группу;

- ИОТ – 10 Инструкция по ОТ при работе с облучателем бактерицидным ОБНП 2(2х15-01) «Генерис»;

- ИОТ – 34 Инструкция по ОТ при работе с ЛВЖ в лабораториях института;

- ИОТ – 72 Инструкция по ОТ при работе с перекисью водорода и органическими перекисными соединениями;

- ИОТ – 86 Инструкция о мерах ПБ в лабораториях;

- ИОТ – 99 Правила работы с микроорганизмами III-IV группы патогенности и возбудителями паразитарных болезней.

- Приложение

INSTRUCTIONS FOR USE OF THE BIOLOG GEN III MICROPLATE™

Intended Use

The GEN III MicroPlate™ test panel provides a standardized micromethod using 94 biochemical tests to

profile and identify a broad range of Gram-negative and Gram-positive bacteria1,2. Biolog’s Microbial

Identification Systems software (e.g.OmniLog® Data Collection) is used to identify the bacterium from its phenotypic pattern in the GEN III MicroPlate.

Description

The Biolog GEN III MicroPlate analyzes a microorganism in 94 phenotypic tests: 71 carbon source

utilization assays and 23 chemical sensitivity assays. The test panel provides Phenotypic Fingerprint of the microorganism that can be used to identify it at the species level.

All necessary nutrients and biochemicals are prefilled and dried into the 96 wells of the MicroPlate. Tetrazolium redox dyes are used to colorimetrically indicate utilization of the carbon sources or resistance to inhibitory chemicals.

The isolate to be identified is grown on agar medium and then suspended in a special “gelling” inoculating fluid (IF) at the recommended cell density. Then the cell suspension is inoculated into the GEN III MicroPlate, 100 μl per well, and the MicroPlate is incubated to allow the phenotypic fingerprint to form. All of the wells start out colorless when inoculated. During incubation there is increased respiration in the wells where cells can utilize a carbon source and/or grow. Increased respiration causes reduction of the tetrazolium redox dye, forming a purple color. Negative wells remain colorless, as does the negative control well (A-1) with no carbon source. There is also a positive control well (A-10) used as a reference for the chemical sensitivity assays in columns 10-12. After incubation, the phenotypic fingerprint of purple wells is compared to Biolog’s extensive species library. If a match is found, a species level identification of the isolate is made.

Determine Appropriate Protocol to Use (Inoculating Fluid and Cell Density)

All protocols are performed in the same manner, the only difference being the choice of inoculating

fluid (IF) and cell density for inoculation.

Protocol A is used for the vast majority of species.

Protocol B is used for a small number of strongly reducing species and capsulated species (primarily some strains of Aeromonas, Vibrio, and spore-forming Gram-positive rods). These species will give a false-positive result in the A-1 well with Protocol A. If this occurs, simply repeat the test using Protocol B.

Protocol C1 is used for slow growing bacteria that typically form pinpoint-sized colonies (less than

1 mm in diameter) on BUG+B Agar in 24 hours of growth. These are primarily microaerophilic and capnophilic Gram-positive cocci and tiny rods. See Table 1. below for a list.

Protocol C2 is used for fastidious, capnophilic, and very oxygen-sensitive bacteria that grow very slowly or not at all on BUG+B Agar. For example, it is used for fastidious Gram-negative species that would most likely be encountered from respiratory tract specimens after cultivation on Chocolate Agar with 6.5% CO2. Some very oxygen-sensitive Gram-positive bacteria also require the higher inoculation density of Protocol C2. See Table 1. below for a list. If unsure of the appropriate test protocol, use protocol A. If the result fails to yield an identification because of a false-positive A-1 well, then use Protocol B. If the result fails because of insufficient positive carbon source reactions, then try, in succession, Protocols C1 and C2.

Table 1. Test Protocols

Protocol IF Cell Density Species

A Nearly all – this is the default protocol

B Strongly reducing and capsule producing GN (e.g., some Aeromonas,

Vibrio) and GP (e.g., some Bacillus, Aneurinibacillus, Brevibacillus,Lysinibacillus, Paenibacillus, and Virgibacillus )

C1 Microaerophilic, capnophilic GP (e.g., Dolosicoccus, Dolosigranulum,

Eremococcus, Gemella, Globicatella, Helcococcus, Ignavigranum,Lactobacillus, Lactococcus, Leuconostoc, Pediococcus, Streptococcus,Weissella, and some Aerococcus, Arcanobacterium, Corynebacterium and Enterococcus sp.)

C2 Fastidious, capnophilic, oxygen sensitive GN (e.g., Actinobacillus, Aggregatibacter, Alysiella, Avibacterium, Bergeriella, Bordetella, Capnocytophaga, Cardiobacterium, CDC Group DF-3, CDC Group EF-4,Conchiformibius, Dysgonomonas, Eikenella, Francisella, Gallibacterium,

Gardnerella, Haemophilus, Histophilus, Kingella, Methylobacterium,Moraxella, Neisseria, Oligella, Ornithobacterium, Pasteurella,Simonsiella, Suttonella, and Taylorella) and GP (Actinomyces,

Aerococcus, Alloiococcus, Arcanobacterium, Carnobacterium, Corynebacterium, Erysipelothrix, Granulicatella, Lactobacillus, Pediococcus, and Tetragenococcus)

TEST PROCEDURE

Preparation

Before starting, prewarm MicroPlates and IF to room temperature and review the entire protocol, including precautions.

Step 1. Culture Organism on Biolog Recommended Agar Media

Isolate a pure culture on Biolog recommended agar media (BUG+B or Chocolate Agar) and incubate at 33° C. Some species may require special culture conditions, for example either lower or higher temperature (26° - 37° C.) and elevated CO2 (6.5% - 10%).

Use of alternative media should be validated. For laboratories that need to use agar media without blood, we recommend using BUG Agar. However, some species will grow extremely slowly or not at all if blood is omitted, for example the genera listed for Protocols C1 and C2 in Table 1. R2A Agar and Tryptic Soy Agar without or with blood (TSA, TSA+B) can be substituted, but they will not culture as wide a range of bacteria as BUG+B. Furthermore, their recipes and performance characteristics from different vendors may vary.

The cells must be freshly grown since many strains lose viability and metabolic vigor in stationary phase. The recommended incubation period for most organisms is 4-24 hours. Sporeforming gram-positive bacteria (Bacillus and related genera) should be grown for less than 16 hours to help minimize sporulation.

If insufficient growth is obtained to inoculate the panel, restreak heavily (as a lawn) onto one or more agar plates. Incubate for 4-48 hours. This should give enough growth to inoculate the panel.

Step 2. Prepare Inoculum

Check the calibration of the turbidimeter periodically. Use an appropriate turbidity standard (85% T or 65% T) and follow instructions in the turbidimeter manual to verify that the turbidimeter is calibrated and operating properly.

Blank the turbidimeter with a clean tube (wiped clean of dirt and fingerprints) containing uninoculated IF. Because the tubes used are not optically uniform, they should be blanked individually. Set the 100% transmittance adjustment knob so that the meter reads 100%.

Prepare the inoculum at the desired turbidity. The target cell density should be in the range

of 90-98%T for Protocols A, B, and C1. Protocol C2 requires a higher cell density of 62-68%T for species that are sensitive to oxygen. Use a cotton-tipped Inoculatorz swab to pick up a 3 mm diameter area of cell growth from the surface of the agar plate.

grasp the swab at its tip and, holding the swab vertically, touch it to the cell growth. For fast growing bacteria, touch a single colony, for medium growing bacteria, touch a cluster of colonies, and for slow growing bacteria touch the first area of confluent growth. Release the bacteria into the IF by rubbing the swab tip against the bottom of the tube containing IF Crush any cell clumps against the tube wall or remove them from the IF by catching them on the swab. Stir the IF with the swab to obtain a uniform cell suspension and read it in the turbidimeter. If the cell density is too low, add more cells. If the cell density is too high, add more IF. For extremely clumpy bacteria that cannot be dispersed directly, use the following procedure. First prepare a dense suspension in 2 ml of IF as follows. Use a sterile wooden Streakerz stick to remove a clump of cell mass from the agar surface without gouging the agar. If the bacteria are extremely dry and embedded in the agar, use the edge of a sterile glass microscope slide to

gently scrape a mass of cells onto the glass slide, again, without gouging the agar. The cells can then be scraped off the glass slide with a sterile Streakerz stick. Then use the Streakerz stick to deposit the cell mass onto the inner wall of a dry tube. Use the Streakerz stick to crush, break up, and spread the clumps of cells against and along the inner wall of the tube. Then add 2 ml of IF, and gradually slide the dispersed cells into the IF. The resulting cell suspension will be a mixture of suspended cells and residual clumps. Stand the tube in a rack for about 5 minutes and allow the clumps to settle to the bottom. Use a small pipet and transfer the suspended cells at the top into a fresh tube of IF to achieve the target cell density.

Step 3. Inoculate MicroPlate

Pour the cell suspension into the multichannel pipet reservoir.

Fasten 8 sterile tips securely onto the 8-Channel Repeating Pipettor and fill the tips by drawing up the cell suspension from the reservoir.

Fill all wells with precisely 100 μl. Be careful not to carry over chemicals or splash from one well into another. The inoculating fluid will form a soft gel shortly after inoculation.

Cover the MicroPlate with its lid and eject the pipettor tips.

Step 4. Incubate MicroPlate

Place the MicroPlate into the OmniLog incubator/reader, or into an incubator, for 3 to 36 hours. Шncubate at 33° C., or use incubation conditions that were found to be optimal for the bacterium in Step 1.

RESULTS

Reading and Interpretation of Results

Read MicroPlates using Biolog’s Microbial Identification Systems software (e.g.OmniLog® Data

Collection). Refer to the User Guide for instructions.

Biolog.’s Microbial Identification Systems Software performs all reading and interpretation of

results.

The color densities in wells of the carbon source utilization assays in columns 1-9 are referenced against the negative control well, A-1. All wells visually resembling the A-1 well should be scored as “negative” (-) and all wells with a noticeable purple color (greater than well A-1) should be scored as “positive” (+). Wells with extremely faint color, or with small purple flecks or clumps are best scored as “borderline” (\). Most species give dark, clearly discernible “positive” reactions. However, it is normal for the “positive” reactions of certain genera to be light or faint purple.

The color densities in wells of the chemical sensitivity assays in columns 10-12 are referenced against the positive control well, A-10. All wells showing significant sensitivity to the inhibitory chemical, with less than half the color of the A-10 well are considered “negative” (-) for growth. All other wells showing normal or near normal purple color (similar to well A-10) are considered “positive” (+). If there is uncertainty about the interpretation, it is best to score the well as “borderline” (\).

“False positive.” color is defined as purple color forming in the negative control well (A-1) and in

other “negative” wells. This is seen with only a few species such as from the genera Aeromonas, Vibrio, and Bacillus. If such a result occurs, the cells are simply retested with Protocol B and IFB.

See Biolog.’s Microbial Identification Systems software User Guide for further assistance

in interpreting identification results.

Precautions

To obtain accurate and reproducible results, the recommendations below must be followed. Read the “Instructions for Use” prior to using the GEN III MicroPlate and follow the procedures. Pure cultures must be used to obtain identifications. The system is not designed to identify individual bacterial strains from within mixed cultures. The most common problem in identification is that microbiologists are not aware that they have a mixed culture. Streaking for isolated colonies may not be sufficient because isolated colonies can arise from a clump of cells as well as a single cell. Bacteria have sticky surfaces and they tightly adhere to other bacteria. This is particularly a problem with mucoid bacteria, fresh environmental isolates, and staphylococci. First, examine cultures with care using a dissecting microscope or some colony magnifying lens, to make sure that only one colony morphology is present in the culture. If no species identification is obtained, you may still have a mixed culture. Restreak the cells onto a multi-chromogenic agar medium and let the original agar plate and the chromogenic agar plate sit at room temperature for 3 or 4 days. Examine both plates carefully, looking for the outgrowth of “bumps” or non-uniform growth in the areas of confluent growth. On the chromogenic agar plate, look for more than one color. If necessary, reisolate the colony types that are present and perform the identification assay a second time.

Culture media and repeated subculturing may affect the results. Strains may produce different phenotypic patterns depending upon how they are cultured prior to inoculation.

Sterile components and aseptic techniques must be used in set-up procedures. Contamination will

affect results.

Disposable glassware should be used to handle all cell suspensions and solutions. Glassware that has been washed may contain trace amounts of soap or detergent that will affect results.

Prewarm the IF and the MicroPlates to room temperature before use. Some species (e.g.,

Neisseria sp.) are very sensitive to cold shocks.

Check the calibration of your turbidimeter carefully and always prepare your inoculum within the specified density range.

Biolog.’s chemistry contains components that are sensitive to temperature and light. Store the inoculating fluids in the dark with refrigeration. Brown or yellow wells in the GEN III MicroPlate indicate deterioration of the chemistry.

Always keep in mind that you are testing the metabolic properties of live cells. Some species can

lose their metabolic vigor when subjected to stresses (e.g., temperature, pH, and osmolarity) for

even a few minutes. To get the best performance possible from these MicroPlates, be aware that

the cells are alive and take care in how you handle them.

Trouble Shooting

If all wells in columns 1-9 are positive, make sure that:

You are using a microorganism that is appropriate for the GEN III MicroPlate. If the bacterium is a

strongly reducing or capsulated species causing false positive color in the A-1 well, repeat the test

using Protocol B and IF-B.

You are not carrying over any nutrients from the agar growth medium into the inoculating fluid.

Your inoculum is free of all clumps.

Your inoculum density is not excessive – check the calibration of your turbidimeter. The A-1 well is not under-filled. It is used as a reference well by Biolog’s Microbial Identification Systems software.

If all wells in columns 1-9 are negative, make sure that:

You are using a microorganism that is appropriate for the GEN III MicroPlate. Oligotrophic species

or extremely slow growing or oxygen sensitive bacteria, for example, may give all negative wells.

Your cells are freshly grown and you have used the recommended agar culture medium.

Your incubation temperature and atmosphere are correct for the organism that is being tested.

The inoculating fluid was stored correctly and was prewarmed prior to use.

You are handling the cells with all disposable hardware (soap residues are toxic).

Your inoculum density is sufficient – check the calibration of your turbidimeter.

The A-1 well is not over-filled. It is used as a reference well by Biolog’s Microbial Identification

Systems software.

Performance Characteristics

The GEN III MicroPlate performance characteristics have been determined by establishing a database using a large collection of microorganisms from diverse sources. The database is designed to give identifications of all species in the database, in accordance with current standards of classical identification methods and current taxonomic nomenclature. To obtain accurate and reproducible results, all procedures and recommendations in these Instructions for Use must be followed precisely.

Limitations

The GEN III MicroPlate is designed to identify pure cultures of Gram-negative and Gram-positive bacteria.

The panel will only identify members of the species in the current database. Other species will usually be

reported out with the message “no identification.” Atypical strains may also yield a low similarity index and therefore will be reported out as “no identification.” This product is not for human in vitro diagnostic use. Some bacterial species are reportable to government and public health agencies in certain

circumstances. For any isolate that is identified as Salmonella or Shigella or E. coli O157:H7, we

recommend confirmation by serology. Neisseria gonorrhoeae identifications should also be confirmed.

Appropriate caution and confirmation should be used for isolates suspected of being Dangerous Pathogens.

Quality Control

Biolog MicroPlates are tested and meet internal quality control standards before being released for sale.

However, some laboratories may desire or may be required to perform independent quality control checks

on each manufacturing lot.

To test the performance of the GEN III MicroPlate use the 2 Gram-negative and 2 Gram-positive strains

specified below using Protocol A. These are available from Biolog as a set (Biolog Catalog No.8050).

1. Escherichia coli ATCC 11775

2. Paenibacillus polymyxa ATCC 842

3. Staphylococcus epidermidis ATCC 12228

4. Stenotrophomonas maltophilia ATCC 13637

Inoculate each bacterium following the TEST PROCEDURE as specified. When lyophilized or frozen

cultures are used, they should be subcultured at least twice before being tested.

Read the panels after appropriate incubation. The resulting identification should correctly correspond to

the identity of the quality control strain. If the identification does not match, review the test procedures and check the purity of your culture. Repeat the test.

Technical Assistance

For help or to report problems with this product contact Biolog Technical Service either by phone (510-

785-2564) by fax (510-782-4639) or by email (tech@biolog.com) during business hours (7:30 A.M. to 5 P.M.

Pacific Standard Time), or contact the Biolog Distribution Partner in your area.

General information, Certificates of Analysis and MSDS can now be found at www.biolog.com.

- Ссылки

1. Bochner, BR 1989. Sleuthing out Bacterial Identities. Nature 339:157-158.

2. Bochner, BR 1989. “Breathprints” at the Microbial Level. ASM News 55:536-539.

3. Biolog, Inc., US Patent # 5,627,045.

Список ознакомления

| № | ФИО | Должность | Дата | Подпись исполнителя | Подпись руководителя |
| --- | --- | --- | --- | --- | --- |
| 1 | 2 | 3 | 4 | 5 | 6 |